## Slide 1: High Flow Oxygen Therapy
Bin Du, MD
Medical Intensive Care Unit
Peking Union Medical College Hospital

## Slide 2: Conflicts of Interest
Air Liquide
Fisher & Paykel
Mindray

## Slide 3: Recommendations for Oxygen Therapy
Cardiac and respiratory arrest
Hypoxaemia (PaO2 < 7.8 kPa, SaO2 < 90%)
Hypotension (systolic blood pressure < 100 mmHg)
Low cardiac output and metabolic acidosis (bicarbonate < 18 mmol/L)
Respiratory distress (respiratory rate > 24/min)
Bateman NT, Leach RM. ABC of oxygen: acute oxygen therapy. BMJ 1998; 317: 798-801

## Slide 4: Low Flow Devices
Oxygen delivery at flow rate below the normal patient inspiratory flow rate of about 30 lpm
Variable FiO2
Pruitt WC, Jacobs M. Breathing lessons: basics of oxygen therapy. Nursing 2003; 33: 43-45

## Slide 5: Conventional Low Flow Devices
6 L.min-1
FiO2 ≈ 0.3 @ 6 lpm
Conventional O2 tubing and nasal cannulas

## Slide 6: Conventional Low Flow Devices
conventional face mask
face mask with reservoir bag
FiO2 ≈ 0.6 @ 10 lpm
FiO2 ≈ 0.8 @ 15 lpm

## Slide 7: High Flow Devices
Oxygen delivery at rates above the normal inspiratory flow rate
Fixed FiO2 maintained
Pruitt WC, Jacobs M. Breathing lessons: basics of oxygen therapy. Nursing 2003; 33: 43-45
FiO2 | Air/O2 ratio | Oxygen flow | Total flow
24 | 25:1 | 3 | 78
28 | 10:1 | 6 | 66
35 | 5:1 | 9 | 54
40 | 3:1 | 12 | 48
60 | 1:1 | 15 | 30

## Slide 8: Humidification of Oxygen Therapy
Isothermic saturation boundary (ISB)
  The point at which inspired air reaches 37°C and 100% RH
  4 to 5 cm below carina
Downward shift of ISB
  Gasping
  High minute ventilation
  Inhalation of dry and cold air
  Artificial airway

## Slide 9: Consequence of Inadequate Humidification

## Slide 10: Consequence of Inadequate Humidification

## Slide 11: High-Flow Nasal Cannula Setup
Nishimura M. High-flow nasal cannula oxygen therapy in adults. J Intensive Care 2015; 3: 15

## Slide 12: High-Flow Nasal Cannula (HFNC)
Potential mechanisms of clinical benefit
  Favorable effects of heated and humidified gas
  High flow rate minimizes entrainment of room air
  Increase in ventilatory efficiency
Potential clinical applications
  Hypoxemic respiratory failure
  Acute cardiogenic pulmonary edema
  Postoperative patients
  Postextubation

## Slide 13: Potential Mechanisms of Clinical Benefit
Favorable effects of heated and humidified gas
  Improved comfort and tolerance
  Facilitating secretion removal and mobilization
  Reduced metabolic cost of breathing
High flow rate minimizes entrainment of room air
Increase in ventilatory efficiency
  Washout of nasopharyngeal dead space
  PEEP effect
  Effects on breathing pattern

## Slide 14: Discomfort During Oxygen Therapy vs. HFNC
Cuquemelle E, Pham T, Papon J, et al. Heated and humidified high-flow oxygen therapy reduces discomfort during hypoxemic respiratory failure. Respir Care 2012; 57: 1571-1577

## Slide 15: Discomfort During Oxygen Therapy vs. HFNC
Cuquemelle E, Pham T, Papon J, et al. Heated and humidified high-flow oxygen therapy reduces discomfort during hypoxemic respiratory failure. Respir Care 2012; 57: 1571-1577

## Slide 16: Dyspnea after Endotracheal Extubation
Rittayamai N, Tscheikuna J, Rujiwit P, et al. High-flow nasal cannula versus conventional oxygen therapy after endotracheal extubation: a randomized crossover physiologic study. Respir Care 2014; 59: 485-590

## Slide 17: Respiratory Rate after Extubation
Rittayamai N, Tscheikuna J, Rujiwit P, et al. High-flow nasal cannula versus conventional oxygen therapy after endotracheal extubation: a randomized crossover physiologic study. Respir Care 2014; 59: 485-590

## Slide 18: Oxygen Therapy after Extubation
Parke R, McGuinness S, Dixon R, et al. Open-label, phase II study of routine high-flow nasal oxygen therapy in cardiac surgical patients. Br J Anaesth 2013; 111: 925-931
 | Usual Care (n = 171) | HFNC (n = 169)
Patients with SpO2/FiO2 ≥ 445 on day 3 | 42.4% (76) | 46.4% (78)
 |  | 0.45
Comfort score (4° after extubation) | 8.1 (2.3) | 7.5 (2.6)
 |  | 0.03
Patients requiring escalation of resp. support | 62.1 (77) | 37.9 (47)
 |  | 0.001

## Slide 19: Effects on Comfort and Tolerance
Greater overall comfort with HFNC in
  Acute respiratory failure mainly due to pneumonia
  Postextubation after cardiac surgery

## Slide 20: Lung Mucociliary Clearance after HFNC
Hasani A, Chapman TH, McCool D, et al. Domiciliary humidification improves lung mucociliary clearance in patients with bronchiectasis. Chron Respir Dis 2008; 5: 81-86

## Slide 21: Metabolic Cost of Gas Conditioning
Dysart K, Miller TL, Wolfson MR, et al. Research in high flow therapy: mechanisms of action. Respir Med 2009; 103:
1400-1405
Etotal = Eg (37 – Tamb) + Evap (44 mg – AHamb)
whereas,
Etotal: the total energy required per liter of inspired gas
Eg: the energy required to raise 1 liter of gas 1°C (roughly 1.2 J)
Evap: the energy required to raise 1 mg of water from 37 to 100°C plus the latent heat of vaporization for 1 mg of water (0.263 and 2.260 J, respectively)
Tamb: the ambient gas temperature
AHamb: the absolute humidity of ambient gas per liter

## Slide 22: Potential Mechanisms of Clinical Benefit
Favorable effects of heated and humidified gas
  Improved comfort and tolerance
  Facilitating secretion removal and mobilization
  Reduced metabolic cost of breathing
High flow rate minimizes entrainment of room air
Increase in ventilatory efficiency
  Washout of nasopharyngeal dead space
  PEEP effect
  Effects on breathing pattern

## Slide 23: FiO2 Reduction after Chest Binding
Hudson mask |  | 
 | 4 lpm | 23.8 (17.4 – 30.3)
 | 12 lpm (humidified ‘single jet’) | 17.8 (8.8 – 26.7)
 | 24 lpm (humidified ‘double jet’) | 12.2 (5.0 – 19.3)
Nonrebreathing mask |  | 
 | 15 lpm | 4.1 (-3.3 – 11.5)
 | 110 lpm | -0.7 (-4.4 – 3.0)
Vapotherm + nasal prongs |  | 
 | 40 lpm (humidified) | -0.3 (-3.1 – 2.5)
Sim MA, Dean P, Kinsella J, et al. Performance of oxygen delivery devices when the breathing pattern of respiratory failure is simulated. Anaesthesia 2008; 63: 938-940

## Slide 24: Performance of Oxygen Delivery Devices
Wagstaff TAJ, Soni N. Performance of six types of oxygen delivery devices at varying respiratory rates . Anaesthesia 2007; 62: 492-503
Hudson mask
Hudson NRM
Nasal cannulae
Thermovent H2 HME
Venturi
Vapotherm

## Slide 25: Performance of Oxygen Delivery Devices
Chart: FiO2
| Nose breating at rest | Mouth breathing at rest | Nose breating with exercise
50 lpm | 0.568 | 0.535 | 0.408
40 lpm | 0.55 | 0.502 | 0.37
30 lpm | 0.519 | 0.486 | 0.321
20 lpm | 0.465 | 0.416 | 0.313
10 lpm | 0.369 | 0.342 | 0.263
Ritchie JE, Williams AB, Gerard C, et al. Evaluation of humidified nasal high-flow oxygen system, using oxygraphy, capnography and measurement of upper airways pressure. Anesth Intensive Care 2011: 39: 1103-1110

## Slide 26: Potential Mechanisms of Clinical Benefit
Favorable effects of heated and humidified gas
  Improved comfort and tolerance
  Facilitating secretion removal and mobilization
  Reduced metabolic cost of breathing
High flow rate minimizes entrainment of room air
Increase in ventilatory efficiency
  Washout of nasopharyngeal dead space
  PEEP effect
  Effects on breathing pattern

## Slide 27: Washout of Nasopharyngeal Dead Space
Ritchie JE, Williams AB, Gerard C, et al. Evaluation of humidified nasal high-flow oxygen system, using oxygraphy, capnography and measurement of upper airways pressure. Anesth Intensive Care 2011: 39: 1103-1110

## Slide 28: PEEP Effect During HFNC: Lung Model Study
Ritchie JE, Williams AB, Gerard C, et al. Evaluation of humidified nasal high-flow oxygen system, using oxygraphy, capnography and measurement of upper airways pressure. Anesth Intensive Care 2011: 39: 1103-1110

## Slide 29: PEEP Effect During HFNC: Lung Model Study
Ritchie JE, Williams AB, Gerard C, et al. Evaluation of humidified nasal high-flow oxygen system, using oxygraphy, capnography and measurement of upper airways pressure. Anesth Intensive Care 2011: 39: 1103-1110

## Slide 30: PEEP Effect During HFNC in Adults
Groves N, Tobin A. High flow nasal oxygen generates positive airway pressure in adult volunteers. Aust Crit Care 2007; 20: 126-131
Spontaneous breathing with HFNC in
Healthy volunteers
Stable COPD
Idiopathic pulmonary fibrosis
Postcardiac surgery

## Slide 31: Nasopharyngeal Pressure During HFNC
Parke RL, Eccleston ML, McGuinness SP. The effects of flow on airway pressure during nasal high-flow oxygen therapy. Respir Care 2011; 56: 1151-1155

## Slide 32: Nasopharyngeal Pressure During HFNC
Parke RL, Eccleston ML, McGuinness SP. The effects of flow on airway pressure during nasal high-flow oxygen therapy. Respir Care 2011; 56: 1151-1155

## Slide 33: Performance of Oxygen Delivery Devices
Ritchie JE, Williams AB, Gerard C, et al. Evaluation of humidified nasal high-flow oxygen system, using oxygraphy, capnography and measurement of upper airways pressure. Anesth Intensive Care 2011: 39: 1103-1110
40 lpm via nose
10 lpm via nose
40 lpm via mouth open

## Slide 34: Breathing Pattern During HFNC
Riera J, Pérez P, Cortés J, et al. Effect of high-flow nasal cannula and body position on end-expiratory lung volume: a cohort study using electrical impedance tomography. Respir Care 2013; 58: 589-596

## Slide 35: HFNC Increases EELV
Corley A, Caruana LR, Barnett AG, et al. Oxygen delivery through high-flow nasal cannulae increase end-expiratory lung volume and reduce respiratory rate in post-cardiac surgical patients. Br J Anaesth 2011; 107: 998-1004
Increased airway pressure was significantly correlated with increased lung volume of 25.6%
End-expiratory lung impedance is a surrogate for end-expiratory lung volume

## Slide 36: Breathing Pattern During HFNC
Corley A, Caruana LR, Barnett AG, et al. Oxygen delivery through high-flow nasal cannulae increase end-expiratory lung volume and reduce respiratory rate in post-cardiac surgical patients. Br J Anaesth 2011; 107: 998-1004

## Slide 37: Physiological Effect of HFNC
Effects on oxygenation
Effects on airway pressure
Effects on lung volume
Effects on CO2 elimination
Appropriate humidification

## Slide 38: FiO2 at Different Respiratory Rate
Wagstaff TAJ, Soni N. Performance of six types of oxygen delivery devices at varying respiratory rates. Anaesthesia 2007; 62: 492-503
2 lpm
6 lpm
10 lpm
15 lpm

## Slide 39: HFNC vs. Conventional Oxygen Therapy
Ricard JD. High flow nasal oxygen in acute respiratory failure. Minverva Anestesiol 2012; 78: 836-841

## Slide 40: Physiological Effect of HFNC
Effects on oxygenation
Effects on airway pressure
Effects on lung volume
Effects on CO2 elimination
Appropriate humidification

## Slide 41: Equation of Motion
Pressure =
Volume
Compliance
+ Flow × Resistance
volume
Flow
Pressure

## Slide 42: End-Expiratory Pressure & Inspiratory Flow
Groves N, Tobin A. High flow nasal oxygen generates positive airway pressure in adult volunteers. Aust Crit Care 2007; 20: 126-131
Open Mouth
Closed Mouth
Increase in airway pressure: 0.8 cmH2O per 10 lpm flow

## Slide 43: End-Expiratory Pressure & Inspiratory Flow
Parke RL, Eccleston ML, McGuinness SP. The effects of flow on airway pressure during nasal high-flow oxygen therapy. Respir Care 2011; 56: 1151-1155

## Slide 44: Physiological Effect of HFNC
Effects on oxygenation
Effects on airway pressure
Effects on lung volume
Effects on CO2 elimination
Appropriate humidification

## Slide 45: HFNC Increases EELV
Corley A, Caruana LR, Barnett AG, et al. Oxygen delivery through high-flow nasal cannulae increase end-expiratory lung volume and reduce respiratory rate in post-cardiac surgical patients. Br J Anaesth 2011; 107: 998-1004
Increased airway pressure was significantly correlated with increased lung volume of 25.6%
End-expiratory lung impedance is a surrogate for end-expiratory lung volume

## Slide 46: EELV Increases with Flow Rate
O2 nasal cannula
NHFO2 30 L.min-1
NHFO2 50 L.min-1
NHFO2 60 L.min-1
20 patients with postcardiac surgery hypoxemia (PaO2/FiO2 < 300) following extubaion
Two groups
NHFO2 (OptiflowTM) flow rate titrated upwards from 30 to 50 lpm
Conventional oxygen therapy
OXIDIA study (NCT 02186)
End-expiratory lung impedance is a surrogate for end-expiratory lung volume

## Slide 47: Physiological Effect of HFNC
Effects on oxygenation
Effects on airway pressure
Effects on lung volume
Effects on CO2 elimination
Appropriate humidification

## Slide 49: HFNC After Endotracheal Extubation
Maggiore SM, Idone FA, Vaschetto R, et al. Nasal high-flow versus Venturi mask oxygen therapy after extubation: Effects on oxygenation, comfort, and clinical outcome. Am J Respir Crit Care Med 2014; 190: 282-288

## Slide 50: Physiological Effect of HFNC
Effects on oxygenation
Effects on airway pressure
Effects on lung volume
Effects on CO2 elimination
Appropriate humidification

## Slide 51: Potential Benefits of HFNC
Spoletini G, Alotaibi M, Blasi F, et al. Heated humidified high-flow nasal oxygen in adults: mechanisms of action and clinical implications. Chest 2015; 148: 253-261
Mechanism | Clinical Benefit
Small, loose-fitting nasal prongs | Enhanced comfort
Heat and humidification | Enhanced comfort
Increased water content of mucus | Facilitated secretion removal Avoidance of desiccation and epithelial injury
Decreased metabolic cost of breathing | Reduced work of breathing
High nasal flow rate | Reduced inspiratory entrainment of room air if mouth closed; more reliable delivery of FiO2
Washout of upper airway dead space | Improved efficiency of ventilation Enhanced oxygen delivery
PEEP | Counterbalance auto-PEEP Decreased work of breathing

## Slide 52: Practical Application of HFNC
Spoletini G, Alotaibi M, Blasi F, et al. Heated humidified high-flow nasal oxygen in adults: mechanisms of action and clinical implications. Chest 2015; 148: 253-261
Flow rate adjustment
FiO2 titration
35 lpm as initial setting due to some PEEP effect and good tolerance
5 – 10 lpm increments if RR fails to drop or breathing remains labored
5 – 10 lpm decrements if not tolerated
Rapid improvement in oxygenation if SpO2 falls substantially below the target level (usually > 90% to 92%)
Further increases in flow rate then used to maintain targeted SpO2 while FiO2 lowered to nontoxic levels (≤ 60%)
Weaning criteria
Flow rate ≤ 35 lpm and FiO2 ≤ 50%

## Slide 53: HFNC in Acute Respiratory Failure
 | Face Mask (n = 20) | HFNC (n = 20)
Dyspnea | 6.8 (4.1 – 7.9) | 3.8 (1.3 – 5.8)
 |  | 0.001
Mouth dryness | 9.5 (8.0 – 10.0) | 5.0 (2.3 – 7.0)
 |  | < 0.001
Overall comfort | 5.0 (2.3 – 6.8) | 9.0 (8.0 – 10.0)
 |  | < 0.001
Roca O, Riera J, Torres F, et al. High-flow oxygen therapy in acute respiratory failure. Respir Care 2010; 55: 408-413

## Slide 54: HFNC in Acute Respiratory Failure
 | Face Mask (n = 20) | HFNC (n = 20)
Total oxygen flow (lpm) | 15.0 (12.0 – 20.0) | 30.0 (21.3 – 38.7)
 |  | < 0.001
Respiratory rate (bpm) | 28 (25 – 32) | 21 (18 – 27)
 |  | < 0.001
PaO2 (mmHg) | 77 (64 – 88) | 127 (83 – 191)
 |  | 0.002
SpO2 (%) | 95 (91 – 97) | 98 (96 – 99)
 |  | 0.002
Roca O, Riera J, Torres F, et al. High-flow oxygen therapy in acute respiratory failure. Respir Care 2010; 55: 408-413

## Slide 55: HFNC vs. Conventional Oxygen Therapy
Ricard JD. High flow nasal oxygen in acute respiratory failure. Minverva Anestesiol 2012; 78: 836-841

## Slide 56: HFNC vs. NIV in Postoperative Hypoxemia
Stéphan F, Barrucand B, Petit P, et al. High-flow nasal oxygen vs noninvasive positive pressure in hypoxemic patients after cardiothoracic surgery: A randomized clinical trial. JAMA 2015; 313: 2331-2339

## Slide 57: HFNC vs. NIV in Postoperative Hypoxemia
Stéphan F, Barrucand B, Petit P, et al. High-flow nasal oxygen vs noninvasive positive pressure in hypoxemic patients after cardiothoracic surgery: A randomized clinical trial. JAMA 2015; 313: 2331-2339
 | BiPAP PS 8/PEEP 4, FiO2 50% (n = 416) | HFNC 50 lpm, FiO2 50% (n = 414)
Treatment failure (%) | 21.9% (91) | 21.0% (87)
 |  | 0.003
ICU Mortality (%) | 5.5% (23) | 6.8% (28)
 |  | 0.66
Skin breakdown (%) | 9.9% (40/403) | 3.2% (13/405)
 |  | < 0.001

## Slide 58: HFNC vs. Face Mask vs. NIV in Hypoxemia
Frat J, Thille AW, Mercat A, et al. High-flow oxygen through nasal cannula in acute hypoxemic respiratory failure. N Engl J Med 2015; 372: 2185-2196

## Slide 59: HFNC vs. Face Mask vs. NIV in Hypoxemia
Frat J, Thille AW, Mercat A, et al. High-flow oxygen through nasal cannula in acute hypoxemic respiratory failure. N Engl J Med 2015; 372: 2185-2196

## Slide 60: HFNC vs. Face Mask vs. NIV in Hypoxemia
Frat J, Thille AW, Mercat A, et al. High-flow oxygen through nasal cannula in acute hypoxemic respiratory failure. N Engl J Med 2015; 372: 2185-2196
 | HFNC (n = 106) | Standard Oxygen (n = 94) | NIV (n = 110)
Intubation at D28 (%) | 38% (40) | 47% (44) | 55% (55)
 |  | 1.45 (0.83 – 2.55) | 1.65 (0.96 – 2.84)
Ventilator-free days | 24 ± 8 | 22 ± 10 | 19 ± 12
 |  | 0.02 | 
ICU Mortality (%) | 11% (12) | 19% (18) | 27% (25)
 |  | 2.55 (1.07 – 6.08) | 2.60 (1.20 – 5.63)
90-day Mortality (%) | 12% (13) | 23% (22) | 28% (31)
 |  | 2.36 (1.18 – 4.70) | 2.33 (1.22 – 4.47)

## Slide 61: HFNC in Acute Respiratory Failure
Improved oxygenation compared with standard oxygen therapy but not as much as NIV
Reduced respiratory rate respiratory rate ≥ 30 bpm after initiation after HFNC predicted the need for intubation
Better tolerance than NIV
Wide application in ICU, high-dependency unit, and emergency department
Spoletini G, Alotaibi M, Blasi F, et al. Heated humidified high-flow nasal oxygen in adults: mechanisms of action and clinical implications. Chest 2015; 148: 253-261

## Slide 62: HFNC in Acute Heart Failure
Carratalá Perales JM, Llorens P, Brouzet B, et al. High-flow therapy via nasal cannula in acute heart failure. Rev Esp Cardiol 2011; 64: 723-725
 | Before HFNC | After HFNC
PaO2 (mmHg) | 73.4 ± 4.3 | 98.8 ± 4.76
 |  | < 0.001
PaCO2 (mmHg) | 53.2 ± 13 | 47.4 ± 8
 |  | 0.109
Respiratory rate (bpm) | 35 ± 2 | 24 ± 3
 |  | 0.002
Heart rate (bpm) | 103 ± 7 | 91 ± 3
 |  | 0.024
Moderate or severe dyspnea | 100% (5) | 20% (1)
 |  | 0.052

## Slide 63: HFNC in Acute CPE
Open flooded alveoli
Improve lung compliance and oxygenation
Enhance cardiac function through the afterload-reducing effect
Spoletini G, Alotaibi M, Blasi F, et al. Heated humidified high-flow nasal oxygen in adults: mechanisms of action and clinical implications. Chest 2015; 148: 253-261

## Slide 64: HFNC in Postoperative Patients
Parke RL, McGuinness SP, Eccleston ML. A preliminary randomized controlled trial to assess effectiveness of nasal high-flow oxygen in intensive care patients. Respir Care 2011; 56: 265-270
 | Nasal high-flow oxygen therapy (n = 30) | Standard high-flow face mask (n = 30)
Protocol accomplishment (%) | 90% (26/29) | 56% (15/27)
 |  | 0.006
Treatment failure (%) | 10% (3/29) | 30% (8/27)
 |  | 0.10
Desaturation (%) | 42% (8/19) | 71% (10/14)
 |  | 0.16
Mean desaturation per patient | 1.86 | 0.79

## Slide 65: HFNC in Postoperative Patients
Postcardiothoracic and vascular surgery
  Lower failure rate (escalation to NIV)
  Fewer desaturations
  Improved thoracoabdominal synchrony
Postcardiac surgery
  Reduced respiratory rate
  Reduced Borg dyspnea score
  Improved or no significant difference in oxygenation
  Increased airway pressure
  Increased EELV by EIT, especially in patients with a higher BMI
  Reduced need for escalation to NIV or intubation
Spoletini G, Alotaibi M, Blasi F, et al. Heated humidified high-flow nasal oxygen in adults: mechanisms of action and clinical implications. Chest 2015; 148: 253-261

## Slide 66: HFNC After Endotracheal Extubation
Rittayamai N, Tscheikuna J, Rujiwit P, et al. High-flow nasal cannula versus conventional oxygen therapy after endotracheal extubation: a randomized crossover physiologic study. Respir Care 2014; 59: 485-590
 | HFNC (n = 9) | NRM (n = 8)
Dyspnea score | 1.6 ± 1.2 | 2.9 ± 1.5
 |  | 0.04
Comfort score | 1.4 ± 0.9 | 1.9 ± 1.1
 |  | 0.07
Respiratory rate (bpm) | 19.8 ± 3.2 | 23.1 ± 4.4
 |  | 0.009
Heart rate (bpm) | 89.5 ± 9.5 | 95.4 ± 10.4
 |  | 0.006

## Slide 67: HFNC After Endotracheal Extubation
Tiruvoipati R, Lewis D, Haji K, et al. High-flow nasal oxygen vs high-flow face mask: a randomized crossover trial in extubated patients. J Crit Care 2010; 25: 463-468

## Slide 68: HFNC After Endotracheal Extubation
Maggiore SM, Idone FA, Vaschetto R, et al. Nasal high-flow versus Venturi mask oxygen therapy after extubation: Effects on oxygenation, comfort, and clinical outcome. Am J Respir Crit Care Med 2014; 190: 282-288

## Slide 69: HFNC After Endotracheal Extubation
Maggiore SM, Idone FA, Vaschetto R, et al. Nasal high-flow versus Venturi mask oxygen therapy after extubation: Effects on oxygenation, comfort, and clinical outcome. Am J Respir Crit Care Med 2014; 190: 282-288

## Slide 70: HFNC After Endotracheal Extubation
Maggiore SM, Idone FA, Vaschetto R, et al. Nasal high-flow versus Venturi mask oxygen therapy after extubation: Effects on oxygenation, comfort, and clinical outcome. Am J Respir Crit Care Med 2014; 190: 282-288

## Slide 71: Potential Clinical Applications of HFNC
Spoletini G, Alotaibi M, Blasi F, et al. Heated humidified high-flow nasal oxygen in adults: mechanisms of action and clinical implications. Chest 2015; 148: 253-261
Application |  | Benefits
Procedures |  | Enhanced oxygenation during endoscopy
Hypoxemic respiratory failure |  | 
 | ARDS | Mild and early
 | Pneumonia | Enhanced oxygenation
 | Idiopathic pulmonary fibrosis | Lower respiratory rate
Cardiogenic pulmonary edema |  | Enhanced oxygenation, reduced dyspnea
Postoperatively |  | 
 | Cardiothoracic and vascular | Improved thoracoabdominal synchrony
 | Cardiac surgery | Increased EELV, less escalation of therapy
Postextubation |  | Improved oxygenation and ventilation, enhanced comfort, less displacement of interface
DNR patients |  | Improved oxygenation and respiratory mechanics

## Slide 72: Predictors of HFNC Failure
Sztrymf B, Messika J, Bertrand F, et al . Beneficial effects of humidified high flow nasal oxygen in critical care patients: a prospective pilot study. Intensive Care Med 2011; 37: 1780-1786

## Slide 73: Predictors of HFNC Failure
 | HFNC Failure (n = 9) | HFNC Success (n = 29)
Respiratory rate (bpm) | 29.1 ± 3.8 | 24.6 ± 5.8
 |  | 0.05
SpO2 (%) | 94.2 ± 7.8 | 97.8 ± 2.6
 |  | 0.0035
Patients with thoraco-abdominal asynchrony (%) | 75% | 10%
 |  | 0.007
Sztrymf B, Messika J, Bertrand F, et al . Beneficial effects of humidified high flow nasal oxygen in critical care patients: a prospective pilot study. Intensive Care Med 2011; 37: 1780-1786

## Slide 74: Late HFNC Failure as Prognostic Factor
Kang BJ, Koh Y, Lim CM, et al. Failure of high-flow nasal cannula therapy may delay intubation and increase mortality. Intensive Care Med 2015; 41: 623-632

## Slide 75: Late HFNC Failure as Prognostic Factor
Kang BJ, Koh Y, Lim CM, et al. Failure of high-flow nasal cannula therapy may delay intubation and increase mortality. Intensive Care Med 2015; 41: 623-632
Conclusions
Failure of HFNC might cause delayed intubation and worse clinical outcomes in patients with respiratory failure. Large prospective and randomized controlled studies on HFNC failure are needed to draw a definitive conclusions.

## Slide 76: HFNC During Endotracheal Intubation
Vourc’h M, Asfar P, Volteau C, et al. High-flow nasal cannula oxygen during endotracheal intubation in hypoxemic patients: a randomized controlled clinical trial. Intensive Care Med 2015; 41: 1538-1548
 | HFNC (n = 62) | HFFM (n = 57)
Nadir SpO2 (%) | 91.5 (80-96) | 89.5 (81-95)
 |  | 0.44
Difficult intubation (%) | 1.6 (1) | 7.1 (4)
 |  | 0.18
IDS score | 3 (2-4) | 3 (2-4)
 |  | 0.63
Ventilator-free days | 14 (0-22) | 5 (0-16)
 |  | 0.09
Compared with HFFM, HFNC as a preoxygenation device did not reduce the lowest level of desaturation.

## Slide 77: HFNC in Postextubation Obese Patients
Corley A, Bull T, Spooner AJ, et al. Direct extubation onto high-flow nasal cannulae post-cardiac surgery versus standard treatment in patients with a BMI ≥ 30: A randomized controlled trial. Intensive Care Med 2015; 41: 887-894

## Slide 78: The Other Side of the Moon
Demoule A, Rello J. High flow oxygen cannula: the other side of the moon. Intensive Care Med 2015; 41: 1673-1675
HFNC should be stopped early in patients who are at high risk of intubation
Further works should aim at defining the subgroups of patients in which HFNC could be beneficial and identifying the respective role and potential benefit of NIV and HFNC for pre-oxygenation during rapid sequence intubation
HFNC should be of benefit to some selected patients rather than to all patients

## Slide 79: Indication of HFNC in Respiratory Failure
Spoletini G, Alotaibi M, Blasi F, et al. Heated humidified high-flow nasal oxygen in adults: mechanisms of action and clinical implications. Chest 2015; 148: 253-261
moderate
mild
severe
severity of hypoxemia
Traditional treatment
Benefits of HFNC
standard high-flow oxygen therapy by mask
low-flow nasal oxygen therapy
NIV
enhanced comfort, more reliable delivery of FiO2, more efficient ventilation
no benefit
more oxygenation and gas flow than necessary
unlikely to reduce WOB despite better tolerance
not enough positive pressure or ventilator assistance

## Slide 80: Recommended Indication of HFNC
Some evidence for a benefit | Children < 24 months of age with viral bronchiolitis Palliative care Acute hypoxemic respiratory failure with rapid improvement Pre-oxygenation before rapid sequence intubation of non-hypoxemic patients Acute cardiogenic pulmonary edema Extubation of patients after pneumonia or trauma Sleep-related hypoventilation in COPD Acute asthma in children
No evidence for a benefit | Extubation of obese patients after cardiac surgery Pre-oxygenation before rapid sequence intubation of hypoxemic patients Flexible bronchoscopy
Some rationale against the use in clinical practice | Acute hypoxemic respiratory failure with late inbutation Bridge to lung transplant
Benefit remains to be investigated | Acute hypoxemic respiratory failure in immunocompromised patients Pre- or inter-hospital transport AECOPD
Demoule A, Rello J. High flow oxygen cannula: the other side of the moon. Intensive Care Med 2015; 41: 1673-1675

## Slide 81: Take Home Message
High-flow nasal cannula as a new method of oxygen therapy
  More comfort and tolerance
  PEEP effect
Wide clinical application in moderate hypoxemia